## Slide 1: theoretical Aspects of patch clamping: Session 2
William A. Coetzee NYU School of Medicine

## Slide 2: Contents
The patch clamp technique: historical background.
Patch clamp configurations.
Errors associated with patch clamping.
Liquid junction potentials

## Slide 3
Em
Command potential
Current
Two-electrode voltage clamping

## Slide 4: Patch clamping: Voltage clamping with a single electrode

## Slide 6
Closed
Open

## Slide 8: Contents
The patch clamp technique: historical background.
Patch clamp configurations.
Errors associated with patch clamping.
Liquid junction potentials

## Slide 9: Patch clamp configurations
Cell-attached (on-cell)
  Loose patch
Whole-cell
  Perforated patch
Inside-out
Outside-out

## Slide 10: The patch clamp amplifier
Offset
Scaling
High-frequency boost
Headstage
+
-
Rpip
Rf
Vout
If
Vpip
Resistive Headstage
I
Vpip
Ipip
Ipip = If = (Vout – Vpip)/Rf

## Slide 11: Measurement of current
Headstage
+
-
Rpip
Rf
Vout
If
Vpip
Ipip
Ipip = If
= Vout - Vpip
+
-
Difference amplifier
Ipip
Vpip
+
-
R
V
V = -Ipip * R
I-V converter

## Slide 12: Measurement of current
Headstage
+
-
Rpip
Rf
Vout
If
Vpip
Ipip
Vpip
+
-
R
V
Difference amplifier
and
I-V converter
Offset

## Slide 13: Offset and gain

## Slide 14: Contents
The patch clamp technique: historical background.
Patch clamp configurations.
Errors associated with patch clamping.
Liquid junction potentials

## Slide 15: Capacitance compensation
What capacitance?
  Pipette capacitance
  Stray capacitance
Why is capacitance bad?
  Limits the bandwidth and speed
  Saturates recording equipment
  Directly affects the patch clamp amplifier feedback
  Contributes significantly to noise
Solution to the problem
  Minimize capacitance
  Compensate capacitance
How is it compensated?
  Electronic circuits within the patch clamp amplifier

## Slide 16: Sources of Capacitance in Patch Clamping
Stray capacitance
  Pipette holder (driven shield)
Capacitance of the immersed tip
  Internal solution separated from bath solution by a non-conductive material (glass)
  Amount of solution inside the pipette
  Depth of the bath solution
  Thickness and type of glass
  Coating (Sylgardtm)

## Slide 17: Capacitance compensation
Headstage
+
-
Rpip
Vpip
Vpip
+
-
V
+
-
Fast and slow components (magnitude and time constant)

## Slide 18: Fast and slow capacitance compensation

## Slide 19: Seal formation and capacitance compensation
1. offset
2. Seal formation
3. Capacitance compensation
5mV

## Slide 20: Seal formation and capacitance compensation
5mV

## Slide 21: Electrical properties of the membrane
Rs
Vpip
Rm
Cm
ΔVpip
Iin
Iss
Specific cell capacitance is ~1µF/cm2

## Slide 22: Series resistance
Rpip
Ra
Rm
Rs
Vpip
Rm
Cm
Vm
Vpip
Vm
Rs = Rpip + Ra + Rref + other
Rref
“Series resistance and space clamp artifacts can render worthless the handsomest results once they are subjected to critical review”

## Slide 23: Series resistance
Rpip
Ra
Rm
Rs
Vpip
Rm
Cm
Vm
Vpip
Vm
Rref
Steady-state errors
Vm ≠ Vpip
Vm = Vpip – (Im x Rs)
Im

## Slide 24: Series resistance
Vm deviates from Vpip
Im is underestimated and kinetically distorted
With Rs
Amstrong & Gilly (Methods in Enzymology)

## Slide 25: Series resistance
Rs
Vpip
Rm
Cm
Vm
Vpip
Vm
Im
Temporal errors
τ ≈ Rs x Cm

## Slide 29: Series resistance
Rs
Vpip
Rm
Cm
Vm
Vpip
Vm
Im
Temporal errors
τ ≈ Rs x Cm
The solution:
Patch small cells (minimize Cm)
Minimize Rs (low Rpip & good break-in)
Supercharge Vcmd

## Slide 30: Series resistance compensation with positive feedback (“correction”)
Headstage
+
-
Rpip
Vpip
Vcmd
+
-
V
+
-
A signal proportional to the measured current is used to increase the command potential
The increased Vcmd compensates for the voltage drop across Rpip
The feedback is positive (noise and instability) (“lag” controls)

## Slide 31: Series resistance compensation by supercharging (“prediction”)
Headstage
+
-
Rpip
Vpip
Vcmd
+
-
V
Brief “charging” pulse
The membrane reaches the final value quicker
Depends on Rs and Cm
No positive feedback (noise, stability)
Does not correct for Rs

## Slide 32: Effects of compensation
Vcmd
Vpip
Vm
Im
No compensation
Cm compensation
Cm and Rs
compensation

## Slide 33: Contents
The patch clamp technique: historical background.
Patch clamp configurations.
Errors associated with patch clamping.
Liquid junction potentials

## Slide 34: Errors associated with patch clamping
Series resistance and membrane capacitance
Space clamping
Leaks
Liquid junction potentials

## Slide 35: Spatial problems
Ri
Rm
∞
-∞
0
Distance (x)
V0
V0/e
Length constant (λ)
λ
The space constant (λ) is around 100μm for cardiac tissue
The value of the space constant depends on ion channel activity (Rm)

## Slide 36: Leak currents
Itotal = Im + IL
= (Vm/Rm) + (Vm/RL)
Rm
RL
Itotal
I
V
EK

## Slide 37: Offset adjustment (liquid junction potential)
1. offset
2. Seal formation
5mV
Is zero really zero?

## Slide 38: Contents
The patch clamp technique: historical background.
Patch clamp configurations.
Errors associated with patch clamping.
Liquid junction potentials

## Slide 39: Liquid junction potential
Cl-
Cl-
Cl-
Cl-
Asp
Asp
K+
K+
K+
K+
K+
K+
Liquid junction potentials are caused by the different mobilities of ions at interfaces between different solutions

## Slide 40: Liquid junction potential
Vpip
Vm = Vpip - VL
Vpip | VL | Vm
-65 mV | -12 mV | -77mV
-70 mV | -2 mV | -72 mV

## Slide 41: Calculation of liquid junction potential
Clampex (v6+)