## Slide 1: theoretical Aspects of patch clamping: Session 3
William A. Coetzee NYU School of Medicine

## Slide 2: Contents
Microelectrodes and reference electrodes
Data acquisition

## Slide 3: Microelectodes
Type of glass
  Borosilicate , alumina silicate, quartz
Pipette geometry
  Wall thickness and length
  Shape
Cleaning the glass and filtering pipette solution
Fire polishing
Coating
Filling (bubbles)

## Slide 4: Reference electrodes
Why needed?
Ag-AgCl wire or pellet
  Fabrication
  Drift
  Effects of changing [Cl-]
“Leaky” bridge electrodes
  Fabrication (agar)
  Insensitivity to solution changes
  Resistance
Ag
AgCl

## Slide 5: Data acquisition
Analog to digital conversion (ADC)
  ADC characteristics
  Triggering
  Clipping
  Bit noise
Digital to analog conversion (DAC)

## Slide 6: The ADC
Number of channels (e.g. 2 to 16)
Full-scale voltage (e.g. -10V to 10V)
Resolution
  8-bit: 28 = 256 points
  12-bit: 212 = 4,096 points
  16-bit: 216 = 65,536 points
Sample rate
  E.g. 1kHz (1000 samples/s)
Example:
  8-bit ADC with -5 to +5 VFS (10V range)
  10,000 mV/256 = 39mV resolution/bit
  Accuracy is ±1 LSB (measurement is off by 39mV)

## Slide 7: Sampling
Bit resolution
Sampling Rate

## Slide 8: Clipping
+10V
-10V

## Slide 9: Aliasing
50Hz
100Hz
250Hz
500Hz
1000Hz
Sampling rate: 1kHz
Sampling rate: 10kHz

## Slide 10: How fast should we sample?
Nyquist criterion: to avoid aliasing, the sampling frequency must be at least 2x larger than the highest frequency contained within the signal
Common practice: sample at a frequency 5x faster than the highest frequency contained in the signal

## Slide 11: Signal conditioning
Filtering
  Passive, active and digital
Amplification and offset

## Slide 12: Filtering
Low-pass filtering

## Slide 13: Filter types
Low-pass filtering
High-pass filtering
Band-pass filtering
Notch filtering (band reject)

## Slide 14: Filter order
Low-pass filtering
1-pole
2-poles
4-poles
8-poles
Poles | Order | Slope
1 | 1st | 6dB/octave
2 | 2nd | 12dB/octave
4 | 3rd | 24dB/octave
8 | 4th | 48dB/octave

## Slide 15: Transfer function
Analog
  Elliptic, Cauer, Chebyshev, Bessel, Butterworth
Digital
  Gausian,
  box-car, etc

## Slide 16: Filtering: time-domain
Low-pass filtering
4-pole Bessel
response
4-pole Butterworth
response

## Slide 17: Lags introduced by the Bessel response
For a perfect signal, the rise time (t10-90) ≈ 0.3/(-3dB) (~300µs at 1kHz)
Oversampling, followed by digital filtering

## Slide 18: Contents
Electricity and bioelectricity
The operational amplifier
A simple voltage-clamp circuit
Patch clamping
  The patch clamp amplifier
  Patch clamp configurations
  Errors associated with voltage clamping
  Microelectrodes
  The reference electrode
Data conditioning, acquisition and analysis
  Signal conditioning
  Data Acquisition
  Curve Fitting

## Slide 19: Curve fitting
Determine that the data are of appropriate type
  Parametric vs non-parametric, etc
Define the goal
  Smoothing, interpolation, parameter determination, modeling, etc.
Choose the model
Parameter constraint
Choose a weighting scheme and initial values
Perform curve fitting
Consider the validity of the data

## Slide 20: Curve fitting example
Concentration (M) | Exp1 | Exp2 | Exp3
1x10-7 | 1 | 4 | 5
3x10-7 | 5 | 4 | 8
1x10-6 | 14 | 14 | 12
3x10-6 | 19 | 14 | 22
1x10-5 | 23 | 24 | 21
3x10-5 | 26 | 24 | 24
1x10-4 | 26 | 25 | 24
3x10-4 | 27 | 31 | 26
1x10-3 | 27 | 29 | 25

## Slide 21
Goal: parameter determination (EC50)
Choose a model:
Curve fitting example
Hill equation:
A1
A2
EC50
Hill

## Slide 22
Parameter constraint
  A1 – yes (A1 = 0)
  A2 – no
  EC50 – no
  Hill – no
No weighting
Curve fitting example
Hill equation:
A1
A2
EC50
Hill

## Slide 23
Choose initial values
  A1 = 0 (fix)
  A2 = 25
  EC50 = 10-5M
  Hill = 1
Curve fitting example
Hill equation:
A1
A2
EC50
Hill

## Slide 24
Perform the fit:
  A1 = 0 (fixed)
  A2 = 25.38
  EC50 = 1.6x10-6M
  Hill = 346199.1
Evaluate parameters
Curve fitting example
Hill equation:

## Slide 25: Evaluate the fit
Is the fitted curve close to the data?
  R2 value
  Runs test
Are the best-fit values scientifically plausible?
How precise are the best-fit parameter values?
  Confidence intervals
Would another model be more appropriate?

## Slide 26: Curve fitting: useful models
Exponential decay:
Macroscopic inactivation
C
0
A0
A0+C
t
A0/e

## Slide 27: Curve fitting: useful models
Second order exponential decay
C
t1
A1+A2+C
t2
A1
A2

## Slide 28: Curve fitting: useful models
Boltzmann equation:
Steady-state inactivation

## Slide 29: Curve fitting: useful models
First order exponential function :
Recovery from inactivation

## Slide 30: Curve fitting: useful models
ATP-sensitivity
Simplified Hill equation:

## Slide 31: Further reading

## Slide 32: Operational amplifers
http://hyperphysics.phy-astr.gsu.edu/hbase/electronic/opampcon.html
Operational Amplifiers.pdf

## Slide 33: An Introduction to Electrophysiology Equipment and The laboratory Setup
Methods in Enzymology, Vol 207. Ed. Rudy & Iverson. Chapter 1: "Patch Clamp Techniques: an Overview" by Michael Cahalan and Erwin Neher
Single Channel Recording, Edited by Sakmann and Neher, Plenum Press, 1995. Chapter 1: Introduction to Patch Clamping by R. Penner

## Slide 34: The Patch Clamp Amplifier
Single Channel Recording, Edited by Sakmann and Neher, Plenum Press, 1995. Chapter 4: “Electronic Design of the Patch Clamp” by F.J. Sigworth

## Slide 35: Errors Associated with Voltage Clamping - Access Resistance, Series Resistance, Space Clamp and Leaks
Methods in Enzymology, Vol 207. Ed. Rudy & Iverson. Chapter 5: "Access Resistance and Space Clamp Problems Associated with Whole-Cell Voltage Clamping" by Amstrong and Gilly
Methods in Enzymology, Vol 207. Ed. Rudy & Iverson. Chapter 6: "Correction of Liquid Junction Potentials in Patch Clamp Experiments" by Neher

## Slide 36: Noise
The Axon Guide, Chapter 12
Sakmann and Neher, Chapter 5

## Slide 37: Microelectrodes
Current Protocols in Neuroscience (1997) Unit 6.3: Fabrication of Patch Pipets. J.L. Rae and R.A. Levis
Methods in Enzymology, Vol 207. Ed. Rudy & Iverson. Chapter 3: "Patch Clamp Electrodes" by Rae and Lewis

## Slide 38: Patch Clamp Techniques Whole-cell Voltage Clamping
Current Protocols in Neuroscience (1997) Unit 6.6: Whole-Cell Voltage Clamp Recording Meyer B. Jackson (ns0606-Whole-Cell Voltage Clamp Recording)

## Slide 39: Patch Clamp Techniques Single Channel Recording
Current Protocols in Neuroscience (1997) Unit 6.8: Single Channel Recording. Meyer B. Jackson (ns0608-Single-Channel Recording.pdf)
Chapter 5, The Plymouth Workshop Handbook
The Axon Guide, Chapter 5: Advanced Methods In Electrophysiology

## Slide 40: Patch Clamp Techniques Rapid Solution Changes
Current Protocols in Neuroscience (1998) Unit 6.9: Rapid Solution Application Methods. Cha-Min Tang
ns0609-Rapid Solution Application Methods.pdf

## Slide 41: Patch Clamp Techniques Perforated Patch Clamping
The Axon Guide, Chapter 5: Advanced Methods In Electrophysiology

## Slide 42: Patch Clamp Techniques Macropatch Recordings
The Axon Guide, Chapter 5: Advanced Methods In Electrophysiology

## Slide 43: Patch Clamp Techniques Giant Excised Patch Recording
Single Channel Recording, Edited by Sakmann and Neher, Plenum Press, 1995. Chapter 13: “The Giant Membrane Patch.” by Donald Hilgemann
The Axon Guide, Chapter 5: Advanced Methods In Electrophysiology

## Slide 44: Patch Clamp Techniques Recording from Xenopus oocytes
Current Protocols in Neuroscience (1998) Unit 6.12: Voltage Clamp Recordings from Xenopus Oocytes. Nathan Dascal ns0612-Voltage Clamp Recordings from Xenopus Oocytes.pdf

## Slide 45: Patch Clamp Techniques Expression of ion channels in heterologous expression systems
Methods in Enzymology, Vol 207. Ed. Rudy & Iverson. Chapters 26-29

## Slide 46: Patch Clamp Techniques Recoding KATP channel Channels
Current Protocols in Pharmacology. (2003). UNIT 11.6. Electrophysiological Analysis of KATP Channels in Mammalian Cells and Xenopus Oocytes. by Char-Chang Shieh and Murali Gopalakrishnan

## Slide 47: Data Conditioning and Acquisition Signal Conditioning
The Axon Guide, Chapter 6: Signal Conditioning and Signal Conditioners

## Slide 48: Data Conditioning and Acquisition Computers in the Laboratory
The Axon Guide, Chapter 8: Laboratory Computer Issues and Considerations

## Slide 49: Data Conditioning and Acquisition Signal Acquisition
The Axon Guide, Chapter 9: Acquisition Hardware

## Slide 50: Data Analysis Practical Curve Fitting
The Axon Guide – chapter 10 (pp. 218-224)

## Slide 51: Data Analysis Analysis of Whole-Cell Currents
Molleman, Chapter 5

## Slide 52: Data Analysis: Ion activities in solutions; Stability constants of ligands

## Slide 53: Data Analysis: Analysis of Single Channel Records
Molleman, Chapter 6
Chapter 6, The Plymouth Workshop Handbook
Chapter 7, The Plymouth Workshop Handbook

## Slide 54: Data Analysis: Separation and Analysis of Ionic Currents
Chapter 3, The Plymouth Workshop Handbook

## Slide 55: Data Analysis: Kinetic Analysis of voltage-clamp records
Chapter 4, The Plymouth Workshop Handbook

## Slide 56: Data Analysis: Kinetic Analysis of voltage-clamp records
Chapter 4, The Plymouth Workshop Handbook

## Slide 57: Optics and Imaging
Intracellular ion measurement with Fluorescent Indicators
  http://www.probes.com/handbook/sections/0001.html
  http://www.biologists.com/Micro/frame12.html
  Chapter 9, Sakmann and Neher
Flash Photolysis of Caged Compounds
Recording membrane potential using voltage-sensitive dyes

## Slide 58: Basic Fluorescence and Confocal Microscopy
Current Protocols in Neuroscience (1997) Unit 2.1: Theory and Application of Fluorescence. Donald Coling and Bechara Kachar (ns0201-Theory and Application of Fluorescence Microscopy.pdf)
Current Protocols in Neuroscience (1997) Unit 2.2: Basic Confocal Microscopyby Carolyn L. Smith (ns0202-Basic Confocal Microscopy.pdf)